{"text_format":"1","workbook":"xlsx","n_sheets":1,"sheets":["Sheet1"],"not_rendered":{"charts":0,"images":0,"macros":1,"external_links":0,"defined_names":0,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"Sheet1","index":0,"n_rows":18,"n_cols":4,"rows_kept":18,"cut":false,"hidden":false,"merged":[]}
[]
["Pirkimas: Infuzinių tirpalų šildymo priemonė"]
["TECHNINĖ SPECIFIKACIJA"]
[]
["Nr.","Pavadinimas","Kiekis"]
[]
["1.1.","Infuzinių tirpalų šildymo priemonė","7","vnt."]
["1.1.1.","Infuzuojamų tirpalų šildytuvas, skirtas infuzijos metu pašildyti kraują ir skysčius"]
["1.1.2.","Šiluma perduodama per infuzinę liniją apgaubiantį daugkartinį šildymo profilį"]
["1.1.3.","Tinkamas naudoti su standartinėmis 4-5 mm diametro infuzinėmis linijomis"]
["1.1.4.","Efektyviai šildo esant ≥ 1500 ml/val. skysčio srautui"]
["1.1.5.","Tikslinė temperatūra ≥ 37°C"]
["1.1.6.","Skaitmeninis ekranas temperatūrų atvaizdavimui"]
["1.1.7.","Garsiniai ir vizualiniai aliarmai"]
["1.1.8.","Temperatūros paklaida - ne daugiau kaip ± 2 °C"]
["1.1.9.","Apsauga nuo perkaitimo"]
["1.1.10.","Tinkamas tvirtinti prie infuzinio stovo ir medicininių bėgelių"]
["1.1.11.","Suderinama su elektros tinklu 230 V, 50 Hz"]