TECHNICAL SPECIFICATION FOR THE PURCHASE OF EQUIPMENT

internal combustion engine test bench

1. GENERAL REQUIREMENTS FOR THE EQUIPMENT PURCHASED

1. The subject of the procurement – an internal combustion engine test bench with a spark-ignition single-cylinder engine and emission analysis equipment (hereinafter referred to as the Equipment) with installation services. The Equipment will be used to study the combustion process, efficiency and ecological indicators of a spark-ignition engine using mixtures of conventional fuels (gasoline) and hydrogen.

2. The Equipment to be supplied by the Supplier shall be in brand new condition, factory-refurbished components are not allowed. The Equipment shall meet the minimum/necessary requirements set out in this Technical Specification as set out in Table 1 (Required technical characteristics/parameters and other requirements).

3. The Tender price shall include the costs related to the preparation of the site/engineering infrastructure design, delivery, installation, connection and functionality testing of the Equipment at the location specified by the Contracting Authority in accordance with the manufacturer's recommendations.

4. The Tender price shall include the costs related to training of at least 5 employees of the Contracting Authority. After delivering, installing and preparing the Equipment, the Supplier shall provide training of not less than 8 academic hours in the use of the Equipment at a time agreed in advance with the Contracting Authority.

5. The time limit for delivery, installation and commissioning of the Equipment shall be no more than 8 months from the date of entry into force of the contract. This time limit shall include the services for the preparation of the site design.

6. The location for the delivery and installation of the Equipment – Saulėtekio al. 11, 10223 Vilnius.

7. The Supplier shall submit the manufacturer's catalogues or descriptions of the technical characteristics of the proposed Equipment (in case the manufacturer's catalogue does not fully reflect the compliance of the proposed Equipment with the requirements of the technical specification) along with the offer (in PDF format), in either Lithuanian or English. It is desirable that in the descriptions of technical characteristics provided in English, the supplier graphically indicates (i.e., clearly marks — with color highlighting, and/or with arrows, and/or underlines) specific sections of the documents where the required technical characteristics are described, stating which point of the technical requirements they correspond to, or providing a reference to the manufacturer's website for the equipment description, including the exact code/model and its parameters.

8. The Equipment shall be installed and assembled at a site prepared by the Contracting Authority, covering an area of approximately 100 m² (10×10 meters). The Contracting Authority shall prepare the site in accordance with a project developed by the Supplier and mutually agreed upon by the Parties (as provided in Clause 8.1 of the Technical Specification) no later than within 5 (five) months from the date of submission of the project to the Contracting Authority. The aforementioned deadline may be extended by no more than 30 (thirty) days due to adverse weather conditions that prevent the performance of the site preparation works.

8.1. Within 1 (one) month from the entry into force of the contract, the Supplier must provide:

- a project comprising equipment drawings and a detailed schedule of work execution stages;

- a site preparation project;

- technical requirements for site preparation;

- requirements for protection zones (if applicable);

- technical requirements for engineering utilities (including electricity and water supply);

- designated points/locations for engineering utilities connections and outlets (electricity, water, and wastewater);

- load specifications applicable to free-standing containers/buildings intended to be placed on the site.

9. All Equipment shall be installed in a closed and lockable room(s)/container(s), protected against environmental exposure and potential vandalism. All essential components – the spark-ignition single-cylinder internal combustion engine, dynamometer, exhaust gas analyzer, fuel and air consumption measurement equipment, and in-cylinder pressure sensors – are designed, supplied, and installed as a unified testing system. This ensures complete compatibility between components, including signal protocols, sensor calibration, and mechanical mounting interfaces. Integrated containerized test benches are 100% tested at the manufacturer’s facilities and delivered ready for on-site deployment. Such a comprehensive solution eliminates potential mismatches between separate modules, reduces installation and setup time, and minimizes technical risks.

10. The Equipment (including its constituent and/or assembled parts) shall be covered by a warranty of not less than 24 (twenty-four) months, commencing from the date of signing the Equipment Handover-Acceptance Act. During the warranty period, upon receipt of the Purchaser’s notification of Equipment malfunction by email, the Supplier shall restore the Equipment’s technical condition at the location where the Equipment is being used. The Supplier shall rectify any defects no later than within 15 (fifteen) business days from the date of receipt of the Purchaser’s notification. If, due to objective reasons, the defect cannot be rectified within the specified term, the Parties shall sign a bilateral act specifying a new justified deadline for defect rectification. During the warranty period, the repair works must be carried out at the location of the Equipment, except in cases where, due to technical reasons, it is not possible to do so. If the repair cannot be performed at the Equipment's location, the Supplier shall transport the Equipment (or its constituent part/component) to and from the repair site at its own cost.

11. The Supplier shall be required to submit documentation evidencing its experience in supplying internal combustion engine, automotive, or transmission test benches, where the energy generated is supplied, via a high-voltage converter, either to a high-voltage battery equipped with a Battery Management System (BMS) or to the electrical grid. The Supplier must also demonstrate experience in implementing integrated control systems intended for automated testing, as well as systems for data and signal acquisition, storage, processing, and visualization.

12. The hydrogen supplied to the Equipment will be at a pressure of 30 ± 5 bar, i.e., between 25 and 35 bar.

Table 1. Required technical characteristics/parameters and other requirements

| Row No. | Meaning | Required minimum technical characteristics | Proposed technical indicators. Specific technical characteristics must be specified [To be completed by the Supplier] |
| --- | --- | --- | --- |
| 1. | Purpose of the equipment | The Equipment will be used to study the combustion process and the efficiency and ecological indicators of a spark-ignition engine that uses mixtures of conventional gasoline and hydrogen. | Manufacturer (indicate): ................. Model (indicate, if any): ..................... Code (indicate, if any): ........................ |
| 2 | Components of the equipment | - Spark-ignition single-cylinder internal combustion engine (hereinafter referred to as ICE) | Manufacturer (indicate): ................. Model (indicate, if any): ..................... Code (indicate, if any): ........................ |
|  |  | - Test bench | Manufacturer (indicate): ................. Model (indicate, if any): ..................... Code (indicate, if any): ........................ |
|  |  | - Exhaust gas analyzer | Manufacturer (indicate): ................. Model (indicate, if any): ..................... Code (indicate, if any): ........................ |
|  |  | - Liquid fuel (petrol, ethanol, buthanol) consumption measurement equipment | Manufacturer (indicate): ................. Model (indicate, if any): ..................... Code (indicate, if any): ........................ |
|  |  | - Gas (hydrogen, propane-butane, biomethane, natural gas) consumption measurement equipment | Manufacturer (indicate): ................. Model (indicate, if any): ..................... Code (indicate, if any): ........................ |
|  |  | - Air flow measurement equipment | Manufacturer (indicate): ................. Model (indicate, if any): ..................... Code (indicate, if any): ........................ |
|  |  | - Cylinder pressure measuring (pressure indicating) equipment | Manufacturer (indicate): ................. Model (indicate, if any): ..................... Code (indicate, if any): ........................ |
|  |  | - Additional equipment | Manufacturer (indicate): ................. Model (indicate, if any): ..................... Code (indicate, if any): ........................ |
|  |  | - Container | Manufacturer (indicate): ................. Model (indicate, if any): ..................... Code (indicate, if any): ........................ |
| 2.1 | Spark-ignition single-cylinder internal combustion engine | - For scientific research; |  |
|  |  | - hydrogen inlet pressure 30 5 bar; |  |
|  |  | - maximum ICE power from 10 to 30 kW; |  |
|  |  | - number of cylinders – 1; |  |
|  |  | - engine displacement not less than 430 cm3 and not more than 600 cm3; |  |
|  |  | - bore from 80 to 90 mm; |  |
|  |  | - stoke from 85 to 95 mm; |  |
|  |  | - 4 valves; |  |
|  |  | - variable intake and exhaust valve timing; |  |
|  |  | - maximum crankshaft speed, not less than 5000 RPM; |  |
|  |  | - variable compression ratio: range at least from 9 to 15; |  |
|  |  | - liquid fuel supply: port injection and direct injection; |  |
|  |  | - hydrogen supply: port injection; |  |
|  |  | - adjustable ignition timing; |  |
|  |  | - engine exhaust system with two parallel branches: one branch with a three-way catalytic converter (TWC) and a gasoline particulate filter (GPF) downstream of it; the other branch without any emission control equipment (simple pipe): - the exhaust gas flow is directed to one or another branch by mechanically or electronically controlled valves (valve); - differential pressure sensor that measures the pressure of the exhaust gas flow before and after the GPF; - measurement of exhaust gas temperature before and after the TWC; - measurement of exhaust gases before and after the catalytic converter; |  |
|  |  | - coolant temperature measurement in the engine block or head; |  |
|  |  | - measuring the temperature of the air sucked into the engine; |  |
|  |  | - oxygen sensor (wideband lambda sensor) before the catalyst; |  |
|  |  | - oxygen sensor (wideband lambda sensor) after the catalyst; |  |
|  |  | - exhaust gas recirculation (EGR) system; |  |
|  |  | - the engine operation is controlled by an Electronic Control Unit (ECU) with an unlimited duration and functionality license. The license permits the use of the ECU for academic, scientific, and commercial purposes; |  |
|  |  | - the ECU shall provide all internal combustion engine operating parameters over the CAN communication network, accessible via the OBD-II diagnostic system. |  |
| 2.2 | Test bench | - Automated engine control system with the ability to maintain the desired engine speed, torque or throttle opening angle; |  |
|  |  | - automated control system for the test bench with data collection and processing equipment (PC, sensors, wiring, control and measurement modules and cabinets); |  |
|  |  | - test bench control and data acquisition software DEWESoft or equivalent. During testing, all measurement data must be collected, stored in the computer memory, and made available for presentation or printing; |  |
|  |  | - control desk/station for the user staff (for PC, keyboard, screen, etc.); |  |
|  |  | - brake/generator with the frame and fixtures for both the ICE and the brake, with the connecting shaft and safety covers; |  |
|  |  | - brake – permanent magnet synchronous motor (hereinafter referred to as PMSM) or similar controlled by appropriate controller integrated with automated control system for the test bench; |  |
|  |  | - at least the following bench operating modes (control modes available for dynamometer): - throttle (ICE) / constant RPM (brake / e-motor); - throttle (ICE) / constant torque (brake / e-motor); - constant torque (ICE) / constant RPM (brake / e-motor); - constant RPM (ICE) / constant torque (brake / e-motor); - throttle (ICE) / torque (~RPM2/ ~RPM3) (brake / e-motor); |  |
|  |  | - control of the throttle of an ICE – both configurable voltage output and mechanical servo motor, integrated with the test bench control system; |  |
|  |  | - control of the dynamometer (brake – PMSM): the test bench control system must include HV inverter to control PMSM motor generating 3-phase power signals in response to PMSM motor integrated encoder signals to ensure requested rotational speed, torque, or energy recuperated at settable power/current/torque; |  |
|  |  | - torquemeter between the ICE and the brake; |  |
|  |  | - control of the Battery Management System (BMS) managing the high-voltage (HV) battery, including the control of the DC/AC inverter to discharge the battery if it becomes fully charged; |  |
|  |  | - automated safety features to monitor test bench parameters like temperatures, pressures, voltages, etc., settable in software; |  |
|  |  | - integration of measurement from all available sensors mounted in the ICE; |  |
|  |  | - cooling system for the ICE and the e-motor (if necessary), with external coolers on the roof of the container (or next to container); |  |
|  |  | - during the tests, the electrical energy generated by the electric motor-brake of the bench must be supplied to a high-voltage DC battery, which will be provided by the Contracting Authority, and which has the following parameters: - AC charging circuit: 3×340-530 VAC, 3×50 Hz, from 25 to 35 kW; - DC charging circuit voltage from 650 to 825 VDC, from 25 to 35 kW; - battery capacity not less than 70 kWh; - maximum current not less than 750 A; - three-phase AC output circuit: 3×400 VAC, 3×50 Hz, not less than 40 kVA; |  |
|  |  | - test bench system to be interfaced/compatible with the HV battery in the scope of the voltage range, instantaneous power/current value to observe/adapt to HV battery parameters provided by battery BMS such as battery state of charge, battery temperature, absolute values of voltage and maximum value of current the battery can safely accept; |  |
|  |  | - test bench control system has to incorporate measurement sensors to measure HV voltages and currents during test bench operations, and their readouts have to be integrated into the control system of the test bench; |  |
|  |  | - the exhaust gas extraction system of the VDV, which does not affect engine performance, i.e., ensuring that the extraction fan does not reduce the pressure in the exhaust pipe or that excessively long piping does not create additional pressure in the VDV exhaust pipe; |  |
|  |  | - installed liquid fuel tank with a capacity of 5 to 10 liters, with piping to the fuel consumption measurement device and to the ICE. |  |
| 2.3 | Exhaust gas analyzer | - Integrated into a common data acquisition and processing system; |  |
|  |  | - measurement component: CO, CO2, THC, CH4, NOx, NO, H2, O2; |  |
|  |  | - measuring range: - CO (low concentrations): not less than 0 to 5000 ppm; - CO (high concentrations): not less than 0 to 5000 ppm; - CO2: not less than 0 to 20 Vol. %; - THC: not less than 0 to 60 000 ppmC; - CH4: not less than 0 to 25 000 ppmC; - NOx: not less than 0 to 10 000 ppm; - NO: not less than 0 to 10 000 ppm; - H2: not less than 0 to 10 000 volppm; - O2: not less than 0 to 25 Vol. %; |  |
|  |  | - repeatability, not more than 0.5 %; |  |
|  |  | - response time not more than 2 s; |  |
|  |  | - no fewer than two measurement points of exhaust gases, e.g., before and after the TWC; |  |
|  |  | - EGR quantity estimation; |  |
|  |  | - evaluation of the lambda (λ) value; |  |
|  |  | - operating temperature: not less than 5 to 40°C. |  |
| 2.4 | Liquid fuel consumption measurement equipment | - Integrated into a common data acquisition and processing system; |  |
|  |  | - measuring range not less than 0.15 to 20 kg/h; |  |
|  |  | - accuracy: no worse than 0.1 % of rate; |  |
|  |  | - operating temperature range not less than 50 to +150°C; |  |
|  |  | - pressure not less than 30 bar; |  |
|  |  | - ambient temperature range not less than 30 to +50°C. |  |
| 2.5 | Gas consumption measurement equipment | - Integrated into a common data acquisition and processing system; |  |
|  |  | - measuring range not less than 0.15 to 20 kg/h; |  |
|  |  | - accuracy: no worse than 0.1 % of rate; |  |
|  |  | - operating temperature range not less than 50 to +150°C; |  |
|  |  | - pressure not less than 300 bar; |  |
|  |  | - ambient temperature range not less than 30 to +50°C. |  |
| 2.6 | Air flow measurement equipment | - Integrated into a common data acquisition and processing system; |  |
|  |  | - measuring error less than 1 % of measured value; |  |
|  |  | - repeatability, not more than 0.25 % of measured value; |  |
|  |  | - response time not more than 12 ms; |  |
|  |  | - maintenance-free; |  |
|  |  | - measuring range not less than 0 to 100 kg/h. |  |
| 2.7 | Cylinder pressure measuring (pressure indicating) equipment | - Integrated into a common data acquisition and processing system; |  |
|  |  | - ambient temperature range not less than 30 to +50°C; |  |
|  |  | - pressure sensor (piezo type) installed in the spark plug; |  |
|  |  | - maximum measured pressure not less than 200 bar; |  |
|  |  | - sensitivity not less than 11 pC/bar; |  |
|  |  | - sensor operating temperature range not less than 30 to +350°C; |  |
|  |  | - calibration certificate (to provide upon delivery of equipment); |  |
|  |  | - special tools for installing the spark plug with a pressure sensor in the engine (if required). |  |
| 2.8 | Additional equipment | - Supercharging equipment that creates a pressure of at least 2 bar in the engine intake manifold. |  |
| 2.9 | Container | - At least two separate rooms: one for the test engine, test stand and auxiliary equipment; the other for the operator. A fireproof and soundproof partition with a door and a window of at least 1 m2 must be installed between the rooms; |  |
|  |  | - window to the outside from the operator's room; |  |
|  |  | - artificial lighting in both rooms; |  |
|  |  | - both rooms must be insulated and equipped with electric or similar heaters and coolers capable of maintaining an internal temperature of no less than +15°C and no more than +35°C; |  |
|  |  | - at least one room is equipped with a sink with hot and cold water mixer and hand drying equipment; |  |
|  |  | - in accordance with the VILNIUS TECH Visual Identity Guidelines, the container must be painted blue with a white VILNIUS TECH logo on both sides of the container (https://vilniustech.lt/universitetas/universiteto-stilius/322464); |  |
|  |  | - if the Equipment requires an air compressor, it must be included in this purchase. |  |
| Quantity, pcs. |  |  | 1 |
| Total price, Eur (excl. Vat.) |  |  |  |

Remarks:

- If the technical specification refers to a specific model or source, a specific process or brand, patent, types, specific origin or manufacture, then in such case the specified wording shall be “or equivalent”.

- The name (model, manufacturer) and technical specifications (characteristics) of the goods offered by the Supplier, the name and description of the services offered (technical specifications, characteristics), or the description of the works may not be considered confidential information (https://vpt.lrv.lt/uploads/vpt/documents/files/mp/konfidenciali_informacija.pdf).