ANNEX NO. 1 "TECHNICAL SPECIFICATION"

I. GENERAL REQUIREMENTS

1.1. The equipment specified in this technical specification is intended for testing the ecological performance of a spark-ignition engine using mixtures of regular gasoline and hydrogen. The equipment is manufactured individually for the contracting authority in accordance with the requirements set out in this technical specification and/or using serially produced components that the manufacturer combines into a single system.

1. 2. The equipment must be delivered to the address specified by the contracting authority in Vilnius, Lithuania, the equipment must be installed, the packaging materials must be removed (disposed of) and the personnel must be trained no later than 7 months from the date of entry into force of the purchase contract. Upon the contracting authority's written request regarding the delayed site installation or other unforeseen circumstances, the Supplier must postpone the installation of the equipment, preparation for work and personnel training for a period specified by the contracting authority not exceeding 6 months.

1.3. The supplier must submit to the contracting authority, within 30 days from the date of entry into force of the purchase contract:

1.3.1. the supplier has filled in the tables in Section III of this technical specification, indicating the equipment to be delivered and its specific technical characteristics.

1.3.2. manufacturer name and model name, if the component of the equipment to be delivered, for which a separate price is indicated in the tender, is not manufactured individually for the contracting authority under this procurement contract, but is a serially produced product. In such a case, the supplier must provide links to the manufacturer’s internet website with product specifications​ and add attache printed material in PDF format.

If there is no information on the manufacturer's website that would confirm the conformity of the product with the requirements of Section III of this technical specification, a written confirmation of the conformity of the characteristics of the offered products according to a specific point of the technical specification (a free-form declaration that the offered product complies with a specific requirement of the procurement technical specification) shall be submitted by an employee of the manufacturer or an employee of the manufacturer's authorized representative.

The contracting authority shall verify the information provided no later than within 15 days and notify the supplier of the results of the assessment. The equipment delivered by the supplier must comply with all the requirements of Chapter III of these technical specifications. If non-compliance of the delivered goods is detected, the supplier must pay the contracting authority a fine equal to 20 percent of the value of the non-compliant goods within 15 days of the contracting authority's request or replace the goods with those that comply with the requirements. Goods that do not comply with the requirements specified in Chapter III of these technical specifications shall not be accepted.

1. 4. The supplier must not submit the information requested in Section III (tables) of these technical specifications together with the offer, as this information will have to be submitted after the entry into force of the procurement contract, as specified in point 1.3.

1.5. Any references in the technical specification or other parts of the procurement documents to specific models or sources of supply, specific processes, specific materials characteristic of the goods supplied or services provided by a specific supplier, or trademarks, patents, types, specific origin or production, are for informational purposes only and the supplier is not obliged to offer and/or use the products of specific manufacturers. Any references in the technical specification or other parts of the procurement documents to specific certificates and standards (if any) may be applied (used) equivalent to those specified.

1. 6. The goods must be brand new. " new ", " renew "/" refurbished " components are not permitted.

1. 7. Personnel training must include training of at least 5 persons for at least 8 academic hours.

1.8. The Goods (including their components/complete parts) must be provided with a warranty of at least 24 (twenty-four) months, which begins to be calculated from the date of signing the Goods Transfer-Acceptance Act. During the warranty period, the Supplier, having received the Buyer's notification of the Goods defect by e-mail, shall restore the technical condition of the Goods at the place of use. The Supplier must eliminate the defects no later than within 15 (fifteen) working days from the date of notification to the Supplier. If, for objective reasons, the defect cannot be eliminated within the specified period, a bilateral act shall be signed between the parties, indicating another reasonable term for eliminating the defect. During the warranty period, repair work on the Goods must be performed at the location of the Goods (except in cases where this is impossible for technical reasons). If it is not possible to perform repair work at the location of the Goods, the Goods (or their components) shall be removed for warranty repair and returned by the Supplier at its own expense.

1.9. Additional requirements applicable to Lot 1 :

1.9.1. Goods are delivered and installed on the site provided by purchasing organizations, which area about 100 m2 (10×10 m), in the container supplied by the Supplier. The Contracting Authority shall equip the site in accordance with the project prepared by the Supplier and mutually agreed upon by the Parties.

1.9.2. Project must to submit Supplier within 30 days from the date of entry into force of the purchase contract, which consists of the following documents:

- project with equipment drawings and works performance in stages ;

- requirements pitches for installation ;

- requirements protection zones ( if is );

- requirements engineering communications , power ( electricity , water );

- engineering communications and conclusions points / locations ( electricity , water , sewage );

- freely built containers / buildings loads to the site .

1.9.3. The Supplier shall design, deliver and install all Goods as a single system for the purpose specified in clause 1.1 of this technical specification. Installation shall be carried out only after the entire system has been tested at the production site.

II. ENVIRONMENTAL REQUIREMENTS

2.1. The procurement is carried out in accordance with the Order No. D1-508 of the Minister of the Environment of the Republic of Lithuania of 28 June 2011 “On the approval of the list of products for which environmental protection criteria are to be applied to public procurement and purchases, the description of the environmental protection criteria and the procedure for applying the environmental protection criteria that contracting authorities and contracting entities must apply when purchasing goods, services or works” (current version). Environmental requirements are specified in Chapter 12 of the Special Conditions of the Contract.

III. SPECIAL REQUIREMENTS

# Lot No. 1. Internal combustion engine test bench

| Row No. | Name | Required technical parameters | Technical parameters of goods to be delivered (to be filled according clause 1.3 after the contract will enter into force) |
| --- | --- | --- | --- |
| 2.1 | Spark-ignition single-cylinder internal combustion engine | - For scientific research; |  |
|  |  | - maximum output power of the ICE not less than 15 kW; |  |
|  |  | - number of cylinders – 1; |  |
|  |  | - bore from 80 to 90 mm; |  |
|  |  | - stoke from 85 to 95 mm; |  |
|  |  | - 4 valves; |  |
|  |  | - variable intake and exhaust valve timing; |  |
|  |  | - maximum crankshaft speed, not less than 5000 RPM; |  |
|  |  | - variable compression ratio: variation range at least from 9 to 15; |  |
|  |  | - liquid fuel supply: port injection and direct injection; |  |
|  |  | - hydrogen supply: port injection; |  |
|  |  | - adjustable ignition timing; |  |
|  |  | - engine exhaust system with two parallel branches: one branch with a three-way catalytic converter (TWC) and a gasoline particulate filter (GPF) downstream of it; the other branch without any emission control equipment (simple pipe): - the exhaust gas flow is directed to one or another branch by mechanically or electronically controlled valves (valve); - differential pressure sensor that measures the pressure of the exhaust gas flow before and after the GPF; - measurement of exhaust gas temperature before and after the TWC; - measurement of exhaust gases before and after the catalytic converter; |  |
|  |  | - coolant temperature measurement in the engine block or head; |  |
|  |  | - measuring the temperature of the air sucked into the engine; |  |
|  |  | - oxygen sensor (wideband lambda sensor) before the TWC; |  |
|  |  | - oxygen sensor (wideband lambda sensor) after the TWC; |  |
|  |  | - exhaust gas recirculation (EGR) system; |  |
|  |  | - the engine operation is controlled by an Electronic Control Unit (ECU) with an unlimited duration and functionality license. The license permits the use of the ECU for academic, scientific, and commercial purposes; |  |
|  |  | - the ECU shall provide all ICE operating parameters over the CAN communication network, accessible via the OBD-II or equivalent diagnostic system; |  |
|  |  | - ICE liquid cooling system. |  |
| 2.2 | Test bench | - Purpose – to load and operate the test ICE; |  |
|  |  | - during testing, all measurement data must be collected, stored in the computer memory, and made available for presentation or printing; |  |
|  |  | - control station for the user; |  |
|  |  | - brake – (dynamometer) with the frame and fixtures for both the ICE and the brake, with the connecting shaft and safety covers; |  |
|  |  | - brake – permanent magnet synchronous motor – generator (hereinafter referred to as PMSM) or other electric unit, controlled by an appropriate controller integrated with the automated test bench control system; |  |
|  |  | - at least the following bench operating modes: - throttle (ICE) / constant RPM (dynamometer); - throttle (ICE) / constant torque (dynamometer); - constant torque (ICE) / constant RPM (dynamometer); - constant RPM (ICE) / constant torque (dynamometer); - throttle (ICE) / torque (~RPM2/ ~RPM3) (dynamometer); |  |
|  |  | - control of the dynamometer: the test bench control system must include HV inverter to control brake generating 3-phase power signals in response to brake integrated encoder signals to ensure requested rotational speed, torque, or energy recuperated at settable power/current/torque; |  |
|  |  | - torquemeter between the ICE and the dynamometer; |  |
|  |  | - the values from the test bench temperature and speed sensors must be displayed on the screen and recorded in memory; |  |
|  |  | - during the tests, the electrical energy generated by the dynamometer of the bench must be supplied to a electrical grid; |  |
|  |  | - test bench control system has to incorporate measurement sensors to measure HV voltages and currents during test bench operations, and their readouts have to be integrated into the control system of the test bench; |  |
|  |  | - the ICE exhaust gas extraction/removal system must not affect engine operation, i.e., the extraction fan must not reduce the pressure in the exhaust pipe, and an excessively long pipeline must not create additional backpressure in the ICE exhaust system. For example, a continuous long exhaust pipe must be avoided – instead, a relatively thin exhaust pipe should be inserted into a much wider ventilation duct with a high-capacity exhaust gas extraction fan that also draws in ambient air; |  |
|  |  | - installed liquid fuel tank with a capacity of 5 to 10 liters. |  |
| 2.3 | Liquid fuel consumption measurement equipment | - The values of the signals output by the equipment must be displayed on the screen and recorded in the test bench computer’s memory; |  |
|  |  | - measuring range not less than 0.15 to 20 kg/h; |  |
|  |  | - accuracy: no worse than 0.1 % of rate; |  |
|  |  | - operating temperature range not less than 40 to +150°C; |  |
|  |  | - pressure not less than 200 bar; |  |
|  |  | - ambient temperature range not less than 30 to +50°C. |  |
| 2.4 | Gas consumption measurement equipment | - The values of the signals output by the equipment must be displayed on the screen and recorded in the test bench computer’s memory; |  |
|  |  | - measuring range not less than 0.15 to 20 kg/h; |  |
|  |  | - accuracy: no worse than 0.1 % of rate; |  |
|  |  | - operating temperature range not less than 40 to +150°C; |  |
|  |  | - pressure not less than 300 bar; |  |
|  |  | - ambient temperature range not less than 30 to +50°C. |  |
| 2.5 | Air flow measurement equipment | - The values of the signals output by the equipment must be displayed on the screen and recorded in the test bench computer’s memory; |  |
|  |  | - measuring error less than 1 % of measured value; |  |
|  |  | - repeatability, not more than 0.25 % of measured value; |  |
|  |  | - response time not more than 12 ms; |  |
|  |  | - measuring range not less than 5 to 100 kg/h. |  |
| 2.6 | Cylinder pressure measuring (pressure indicating) equipment | - The values of the signals output by the equipment must be displayed on the screen and recorded in the test bench computer’s memory; |  |
|  |  | - ambient temperature range not less than 30 to +50°C; |  |
|  |  | - piezo type pressure sensor; |  |
|  |  | - maximum measured pressure not less than 30 bar; |  |
|  |  | - sensitivity not less than 11 pC/bar; |  |
|  |  | - sensor operating temperature range not less than 20 to +350°C; |  |
|  |  | - calibration certificate (to provide upon delivery of equipment). |  |
| 2.7 | Supercharging equipment | - Supercharging equipment that creates a pressure of at least 2 bar in the engine intake manifold. |  |
| 2.8 | Container | - At least two separate rooms: one for the test engine, test stand and auxiliary equipment; the other for the operator. A fireproof and soundproof partition with a door and a window of at least 1 m2 must be installed between the rooms; |  |
|  |  | - window to the outside from the operator's room; |  |
|  |  | - artificial lighting in both rooms; |  |
|  |  | - both rooms must be insulated and equipped with electric or similar heaters and coolers capable of maintaining an internal temperature of no less than +15°C and no more than +35°C; |  |
|  |  | - at least one room is equipped with a sink with hot and cold water mixer and hand drying equipment; |  |
|  |  | - in accordance with the VILNIUS TECH Visual Identity Guidelines, the container must be painted blue with a white VILNIUS TECH logo on both sides of the container (https://vilniustech.lt/universitetas/universiteto-stilius/322464); |  |
|  |  | - if the Equipment requires an air compressor, it must be included in the delivered system. |  |

# Lot No. 2. Equipment for analyzing emissions from internal combustion engines

| Row No. | Meaning | Required technical parameters | Technical parameters of goods to be delivered (to be filled according clause 1.3 after the contract will enter into force) |
| --- | --- | --- | --- |
| 2 | Exhaust gas (emissions) analysis equipment | - Possibility to export and save measurement data; |  |
|  |  | - measurement component: CO, CO2, THC, CH4, NOx, NO, H2, O2; |  |
|  |  | - measurement principles: - CO/CO₂ – Non-Dispersive Infrared NDIR or Infrared Detector IRD; - NOx – Chemiluminescent detector CLD; - THC/CH₄ – Flame ionization FID; - O₂ – Magneto-pneumatic detection MPD or Paramagnetic detector PMD; - H2 – Mass spectrometry; |  |
|  |  | - Measurement range: - CO (low concentration): not less than 0 to 5000 ppm; - CO (high concentration): not less than 0 to 10% (by volume); - CO2: not less than 0 to 20% (by volume); - THC: not less than 0 to 60,000 ppmC; - CH4: not less than 0 to 20,000 ppmC; - NOx: not less than 0 to 10,000 ppm; - NO: not less than 0 to 10,000 ppm; - O2: not less than 0 to 25% (by volume); - H2: not less than 0 to 10,000 ppm (by volume); |  |
|  |  | - heated measurement lines (hoses) or another method of moisture removal; |  |
|  |  | - result repeatability no greater than 0.5% or span (concentration) drift no greater than 3% over 24 hours; |  |
|  |  | - response time not greater than 2 seconds; |  |
|  |  | - emissions must be measured at no fewer than two points in the internal combustion engine exhaust system. (Note: not applicable to H2 measurement); |  |
|  |  | - operating temperature range: not less than +10 to +35°C. |  |