Annex 1

to the Procurement Conditions

TECHNICAL SPECIFICATION OF A CHEMICAL DETECTOR

2024		No.

- Chemical detector (hereinafter abbreviated as a Detector) shall detect and identify chemical warfare agents (CWA) and toxic industrial chemicals (hereinafter abbreviated as the substances).

- The detector shall be of a portable type.

- Weight of the detector, ready for an operation, shall be not more than 5 kg.

- The detector shall be adapted for the operator, who wears CBRN Individual Protective equipment (IPE).

- The Detector shall show results on the display or control panel with a function switching the display’s or control panel’s backlight on and off.

- The IP rating of the Detector for ingress by dust and liquid according to EN 60529 or equivalent standard shall be at least IP65.

- The detector shall be suitable for use in the temperature range at least from -20°C to +40°C.

- There shall be no GLONASS (or equivalent) receiver (device) built in the detector.

- The detector shall provide an audible and (or) light alarm in the event of chemical detection and identification with possibility to switch it off.

- The detector shall operate on standard commercial available or multi-use custom-made rechargeable batteries.

- The detector shall indicate the battery discharge visually and (or) audibly; there shall be a possibility to switch the indication off.

- The detector shall be compatible to be connected to 220 V and 12 V power supply, if it is foreseen in the construction of the detector.

- Batteries shall ensure power supply for not less than 6 hours of continuous operation;

- The detector shall perform automatic self and calibration tests of all functions and report all errors on the display or control panel.

- The detector shall detect, identify, and show on the screen not less than these substances: sarin (GB), tabun (GA); soman (GD); VX; iprite (HD); lewisite (L); nitrogen mustard (HN), hydrogen cyanide (AC), cyanogen chloride (CK), phosgene (CG), nitric acid (HNO3), sulphur dioxide (SO2);

- The detector must detect sarin (GB), when the concentration of the CWA in the environment is equal or higher than 0,2 mg/m3, and iprite (HD) is equal or higher than 1,0 mg/m3 concentration;

- The detector shall operate in search mode and detect substances no later than in 20 seconds.

- The detector shall identify substances no later than in 3 minutes.

- The detector shall be compatible to transfer contamination measurement data to a computer.

- The detector shall perform self-clean procedure, during search of substances in contaminated area.

- The colour of the external surfaces of the main Detector and all the accessories (except display and connector sockets) shall be painted with either black (RAL 9005 or equivalent), or dark grey (RAL 7013, 7016, 7021, 7022, 7026 or equivalent), or dark green (RAL 6007 „Rifle green“, or equivalent) mat paint.

- Each Detector set shall consist of:

  - The main Detector;

  - The multi-use means of transport made of a strong material and suitable for field (outdoor) use (eg a suitcase, plastic or metal box or equivalent means of transport that cannot be made of paper or cardboard);

  - A battery;

  - A backup battery;

  - A battery charger;

  - A data transfer cable from the detector to a computer;

  - An external data storage (e.g. SD card), if it is foreseen in the construction of the detector.

- The detector shall have not less than two sets of consumables.

- The Detector shall not have any built-in radioactive source and if there is a radioactive source, it shall be of the exemption level.

- The Detector shall be new and unused.

- The Detector shall have a warranty period of at least 24 months. The warranty period starts from the date of signing the deed of acceptance of the buyer and the seller. The seller shall commit to eliminate the defects during the warranty period within a maximum period of 30 calendar days or to replace the defective components with functional ones. During the warranty period, the supplier shall carry out the repair at his own expense, including transportation costs.

- The manufacturer or supplier shall organize training concerning the use of the Detector in Lithuania no later than within three weeks from the delivery of the goods, as well as to issue training certificates to the participants of the training (up to 20 participants).

- The Detector shall have the CE marking (the copy of certification is to be submitted with the tender documents.

- The Detector shall not be a unitary (experimental) model.

- Together with the detector manufacturer will provide manual and other documentation including technical and maintenance instructions, in English language (soft and hard copy).

___________________________________________