Appendix

to the Annex 8 „Proposal form“

(LOT 1)

TABLE OF CONFORMITY OF THE PROPOSED CHEMICAL DETECTOR TO THE TECHNICAL SPECIFICATION

|  | Requirements | Compliance with requirements (to be filled in by the Supplier, indicating the specific meaning (if requested) or whether the offered product complies or does not comply with a specific mandatory requirement) | Document confirming compliance (to be filled in by the Supplier, indicating the names of the documents and/or links to their points proving the compliance of the offered goods with the technical requirements)* |
| --- | --- | --- | --- |
| 1 | 2 | 3 | 4 |
| 1. | Chemical detector (hereinafter abbreviated as a Detector) shall detect and identify chemical warfare agents (CWA) and toxic industrial chemicals (hereinafter abbreviated as the substances). | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 2. | The detector shall be of a portable type. | YES/NO | Link to the source of information |
| 3. | Weight of the detector, ready for an operation, shall be not more than 5 kg. | To specify exact wight of the Detector, ready for an operation ___ kg. | Link to the source of information |
| 4. | The detector shall be adapted for the operator, who wears CBRN Individual Protective equipment (IPE). | YES/NO | Link to the source of information |
| 5. | The Detector shall show results on the display or control panel with a function switching the display’s or control panel’s backlight on and off. | YES/NO | Link to the source of information |
| 6. | The IP rating of the Detector for ingress by dust and liquid according to EN 60529 or equivalent standard shall be at least IP65. | To specify theIP rating of the Detector for ingress by dust and liquid according to EN 60529 or equivalent standard IP__ | Link to the source of information |
| 7. | The detector shall be suitable for use in the temperature range at least from -20°C to +40°C. | To specify exact diapason of temperature the detector is suitable for use: from -__°C to +__°C. | Link to the source of information |
| 8. | There shall be no GLONASS (or equivalent) receiver (device) built in the detector | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 9. | The detector shall provide an audible and (or) light alarm in the event of chemical detection and identification with possibility to switch it off. | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 10. | The detector shall operate on standard commercial available or multi-use custom-made rechargeable batteries. | YES/NO | Link to the source of information |
| 11. | The detector shall indicate the battery discharge visually and (or) audibly; there shall be a possibility to switch the indication off. | YES/NO | Link to the source of information |
| 12. | The detector shall be compatible to be connected to 220 V and 12 V power supply, if it is foreseen in the construction of the detector. | YES/NO | Link to the source of information |
| 13. | Batteries shall ensure power supply for not less than 6 hours of continuous operation; | To specify exact duration of continious work of the batteries __ h. | Link to the source of information |
| 14. | The detector shall perform automatic self and calibration tests of all functions and report all errors on the display or control panel. | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 15. | The detector shall detect, identify, and show on the screen not less than these substances: sarin (GB), tabun (GA); soman (GD); VX; iprite (HD); lewisite (L); nitrogen mustard (HN), hydrogen cyanide (AC), cyanogen chloride (CK), phosgene (CG), nitric acid (HNO3), sulphur dioxide (SO2); | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 16. | The detector must detect sarin (GB), when the concentration of the CWA in the environment is equal or higher than 0,2 mg/m3, and iprite (HD) is equal or higher than 1,0 mg/m3 concentration; | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 17. | The detector shall operate in search mode and detect substances no later than in 20 seconds. | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 18. | The detector shall identify substances no later than in 3 minutes. | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 19. | The detector shall be compatible to transfer contamination measurement data to a computer. | YES/NO | Link to the source of information |
| 20. | The detector shall perform self-clean procedure, during search of substances in contaminated area. | YES/NO | Link to the source of information |
| 21. | The colour of the external surfaces of the main Detector and all the accessories (except display and connector sockets) shall be painted with either black (RAL 9005 or equivalent), or dark grey (RAL 7013, 7016, 7021, 7022, 7026 or equivalent), or dark green (RAL 6007 „Rifle green“, or equivalent) mat paint. | To specify the colour of the external surfaces of the main Detector and all the accessories RAL ____ | Link to the source of information |
| 22. | Each Detector set shall consist of: |  |  |
| 22.1. | The main Detector; | YES/NO | Link to the source of information |
| 22.2. | The multi-use means of transport made of a strong material and suitable for field (outdoor) use (eg a suitcase, plastic or metal box or equivalent means of transport that cannot be made of paper or cardboard); | YES/NO | Link to the source of information |
| 22.3. | A battery; | YES/NO | Link to the source of information |
| 22.4. | A backup battery; | YES/NO | Link to the source of information |
| 22.5. | A battery charger; | YES/NO | Link to the source of information |
| 22.6. | A data transfer cable from the detector to a computer; | YES/NO | Link to the source of information |
| 22.7. | An external data storage (e.g. SD card), if it is foreseen in the construction of the detector. | YES/NO | Link to the source of information |
| 23. | The detector shall have not less than two sets of consumables. | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 24. | The Detector shall not have any built-in radioactive source and if there is a radioactive source, it shall be of the exemption level. | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 25. | The Detector shall be new and unused. | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 26. | The Detector shall have a warranty period of at least 24 months. The warranty period starts from the date of signing the deed of acceptance of the buyer and the seller. The seller shall commit to eliminate the defects during the warranty period within a maximum period of 30 calendar days or to replace the defective components with functional ones. During the warranty period, the supplier shall carry out the repair at his own expense, including transportation costs. | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 27. | The manufacturer or supplier shall organize training concerning the use of the Detector in Lithuania no later than within three weeks from the delivery of the goods, as well as to issue training certificates to the participants of the training (up to 20 participants). | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 28. | The Detector shall have the CE marking (the copy of certification is to be submitted with the tender documents. | YES/NO | Copy of CE Certificate |
| 29. | The Detector shall not be a unitary (experimental) model. | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |
| 30. | Together with the detector manufacturer will provide manual and other documentation including technical and maintenance instructions, in English language (soft and hard copy). | Manufacturer‘s confirmation DO CONFIRM/DO NOT CONFIRM | Not to be completed when the offer is submitted by the manufacturer of the goods. If the offer is submitted by a supplier who is not the manufacturer of the goods, he must submit a declaration from the manufacturer of the goods regarding compliance with the requirements of this point. |

*Notes:

1) Important: The Supplier, when indicating the specific meaning of the goods or their conformity to it, must indicate to the contracting authority which document submitted by the Supplier contains the data confirming/proving the requirement (indicate the specific folder/file in which the document is submitted, the name of the document and the page);

2) each point of the technical requirements that requests a specific value must be based on the manufacturer's technical documentation;

3) each point of the technical requirements requiring manufacturer's approval, if the offer is submitted by the manufacturer of the goods himself, this is confirmed in column 3 and no supporting documents need to be submitted, and the corresponding column in column 4 is not filled in. If the offer is not submitted by the manufacturer of the goods, the supplier must:

a) fill in the corresponding column in column 4;

b) submit a declaration signed by the manufacturer together with the offer confirming compliance with the specific requirement;

4) "or equivalent" is written next to all specified specific materials and/or specific product names. A supplier offering a product with equivalent properties must, at the time of submitting the offer, prove by reliable means that the offered product is equivalent and fully complies with the requirements set out in the technical specification.