Annex 1 to the Specific Conditions

TECHNICAL SPECIFICATION

1. TERMS AND ABBREVIATIONS / GENERAL INFORMATION

1.1.  Buyer / Purchasing Organization / VU – Vilnius University.

1.2. Supplier – a business entity – a natural person, a private or public legal entity, another organization and its division, or a group of such persons, including temporary associations of business entities, with whom the Buyer will enter into the Purchase Agreement.
1.3. Contract – the Purchase Agreement concluded between the Supplier and the Buyer regarding the subject of this Purchase.

1.4. Project - Vilnius university, aiming to implement the project: LMT, P-IRA-25-29 (Contract No. S-IRA-25-10), plans to acquire the item specified below.

2. SUBJECT OF THE PURCHASE

2.1. Purchase Subject – portable electroencephalography system (hereinafter referred to as the product (-s).
2.2. The purchase subject is not divided into parts, so the Supplier must submit an offer for the entire scope of the purchase object as outlined below.

2.3. Delivery Location – VU Faculty of Medicine, Žaliųjų Ežerų 2, LT-08406, Vilnius.

1 Table

| No. | Product name | Quantity and Unit of Measurement | Order Submission |  | Product Delivery / Supply Deadline from Contract Effective Date (calendar days) |
| --- | --- | --- | --- | --- | --- |
|  |  |  | Yes (mark if orders will be submitted on demand, periodically, etc.)** | No (mark if the entire purchased quantity will be delivered at the specified time)* |  |
| 1. | Portable electroencephalography system | 1 set | ☐ | ☒ | 30 c.d. |

2.4. The quantity and (or) amount of products indicated in the above table is final and will not change during the execution of the Contract.

2.5. Order Submission Procedure:

2.5.1. Orders will not be submitted during the term of the Contract. The products must be delivered no later than the deadline specified in Table 1.

3. REQUIREMENTS FOR THE PRODUCTS

3.1. If specific models, sources, specific processes, brand names, patents, types, specific origin, or production are mentioned in the purchase documents, they may be replaced by equivalent alternatives.

2 Table

| No. | Required parameter** | Required parameter value | Offered parameter value (to be filled by the Supplier) |
| --- | --- | --- | --- |
|  | Manufacturer, model | Specify the model and manufacturer. It is mandatory to provide a link to the manufacturer’s (or an equivalent) website or a copy of the technical documentation containing information about the characteristics of the proposed product. If the Supplier is the manufacturer of the product, their own declaration is acceptable. |  |
|  | Electroencephalogram (EEG) collection type | Wireless |  |
|  | Number of channels | Not less than 64 |  |
|  | Channel configuration | At least 8 bipolar channels configurable for measuring external (non-brain) activity |  |
|  | Data transfer method | Bluetooth (5 or newer) or equivalent |  |
|  | Sampling rate | At least 4kHz |  |
|  | Data resolution | At least 24 bit |  |
|  | Referred noise | Not more than 1μV |  |
|  | Motion tracking | 3D head motion tracking |  |
|  | Battery* | Integrated |  |
|  | Battery working capacity | At least 10 hours |  |
|  | Battery charging connector* | Wired and wireless |  |
|  | Memory | Integrated, possibility of recording to internal memory |  |
|  | Weight (amplifier)* | Not more than 80 gr |  |
|  | Functionalities | - Possibility to simultaneously record data from two or more subjects - Possibility to simultaneously record video material synchronized with EEG - Mounting the amplifier on the head, at the back of the head - TTL (Transistor–Transistor Logic, hardware based) and LSL (Lab Streaming Layer, software based) - Automatic algorithm for real-time artefact removal - Data collection device with specialized data collection and visualization software |  |
|  | Adds-on* | 1. Gel/saline based EEG caps. Three sizes 1x54 (S); 1x56 (M); 1x58(L). 2. Preparatory supplies – EEG gel (not less than 1kg), electrode cleaning supplies |  |
|  | Warranty* | Not less than 12 months. The product warranty must cover free replacement of defective parts, repair work, fault elimination, or replacement of the entire product to ensure that the equipment operates fully throughout the warranty period. |  |

**Provide the technical specifications of the offered equipment, excluding those marked with *, along with reliable supporting documents (e.g., manufacturer's product description or a link to the manufacturer's website, or othel equivalent documents).

An equivalent purchase object is considered to be one whose properties are not inferior (i.e., the same or better) compared to the requirements specified for the purchased object in the purchase documents, and the proposed equivalent purchase object can be used for its intended purpose without any restrictions (including but not limited to the following):

No additional modifications to interacting components are required.

Its use will not result in faster wear, failures, and/or loss of warranty for interacting components.

The expected service life is not shorter.

It is not inferior in terms of technical advancement.

When proposing an equivalent purchase object, it is mandatory to provide documents proving compliance with the requirements set for the purchase object. Such documents could include a research report or certificate issued by a conformity assessment body established in the Republic of Lithuania, as well as certificates issued by equivalent conformity assessment bodies established in other countries. If the Supplier is unable to obtain the specified certificates or research reports due to circumstances beyond their control, and provides objective written evidence proving that the proposed equivalent purchase object meets the requirements or criteria specified in the Technical Specification, the evaluation criteria for the offer, or the terms of the purchase agreement, the Buyer will accept other suitable measures. However, the Supplier’s self-declaration without specific, technical evidence will not be considered an acceptable measure. The Buyer reserves the right to conduct a risk assessment of danger if the documents submitted to prove the equivalence of the proposed goods are insufficient