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	1.
	
Laba diena,

Dėkojame už pateiktus atsakymus. Išanalizavę pirkimo techninės specifikacijos reikalavimus, atkreipiame dėmesį į sąlygą dėl įrangos pritaikymo H2S koncentracijai iki 7 mg/m³.

Norime atkreipti jūsų dėmesį, kad absoliuti dauguma Europos gamintojų savo įrangą testuoja, sertifikuoja ir bazinę garantiją teikia pagal bendraeuropinį dujų kokybės standartą EN 16726, kuriame nustatyta 5 mg/m³ riba. Nors teoriškai LR Energetikos ministro įsakymas Nr. 1-194 leidžia H2S koncentraciją iki 7 mg/m³, praktikoje ir Lietuvos gamtinių dujų perdavimo tinkle ši vertė yra artima nuliui, todėl įrangos eksploatacija nedaro jokios įtakos nei matavimų tikslumui, nei įrangos ilgaamžiškumui.

Griežtas ir nelankstus reikalavimas atitikti 7 mg/m³ parametrą be galimybės pagrįsti įrangos tinkamumą kitais būdais, dirbtinai riboja konkurenciją ir eliminuoja aukščiausios kokybės Europos gamintojus, kurių standartinė specifikacija remiasi EN 16726.

Siekiant užtikrinti sąžiningą konkurenciją ir gauti geriausius techninius pasiūlymus, prašome jūsų patikslinti techninę specifikaciją (arba patvirtinti), kad:
Pasiūlymas bus laikomas tinkamu ir neatmestinas, jeigu tiekėjas pasiūlys įrangą, kurios bazinis sertifikavimas atitinka EN 16726 (iki 5 mg/m³), tačiau kartu pateiks oficialų įrangos gamintojo raštą, patvirtinantį, kad įranga yra tinkama eksploatuoti Lietuvos dujų tinkle prie H2S koncentracijos iki 7 mg/m³ (remiantis įsakymu Nr. 1-194), ir kad šis nedidelis skirtumas nepažeidžia techninių parametrų bei visiškai išlaiko gamintojo garantiją.

Prašome įvertinti šį prašymą ir taip pat atsižvelgti į realią šių dienų dujų sudėtį. Galbūt galite pateikti vidutinį šių metų H2S kiekį dujose?

Atsižvelgiant į tai, kad tokio oficialaus gamintojo rašto parengimas ir derinimas užima papildomo laiko, taip pat labai prašome pratęsti pasiūlymų pateikimo terminą bent iki Liepos 3 d. Tai suteiktų galimybę tinkamai pasiruošti ir pateikti jums maksimaliai konkurencingą ir visus reikalavimus atitinkantį pasiūlymą.

Prašome įvertinti šį prašymą ir lauksime jūsų komentarų.
	Good day,

Thank you for the answers provided. After analysing the requirements of the technical specification, we would like to draw your attention to the condition regarding the equipment’s suitability for H₂S concentrations up to 7 mg/m³.

We would like to note that the vast majority of European manufacturers test, certify, and provide their standard warranty based on the pan‑European gas quality standard EN 16726, which sets a limit of 5 mg/m³. Although, in theory, the Order No. 1‑194 of the Minister of Energy of the Republic of Lithuania allows H₂S concentrations up to 7 mg/m³, in practice—and specifically within the Lithuanian natural gas transmission network—this value is close to zero. Therefore, the operation of the equipment has no impact on measurement accuracy or equipment longevity.

A strict and inflexible requirement to meet the 7 mg/m³ parameter, without allowing the suitability of the equipment to be justified in other ways, artificially restricts competition and excludes top‑quality European manufacturers whose standard specifications are based on EN 16726.

To ensure fair competition and to receive the best technical proposals, we kindly ask you to clarify the technical specification (or confirm) that:
A proposal will be considered compliant and will not be rejected if the supplier offers equipment whose basic certification meets EN 16726 (up to 5 mg/m³), but also provides an official letter from the manufacturer confirming that the equipment is suitable for operation in the Lithuanian gas network at H₂S concentrations up to 7 mg/m³ (in accordance with Order No. 1‑194), and that this small difference does not affect the technical parameters and fully preserves the manufacturer’s warranty.

We kindly ask you to consider this request and also take into account the actual gas composition today. Could you perhaps provide the average H₂S concentration in the gas this year?

Considering that preparing and coordinating such an official manufacturer’s letter requires additional time, we also kindly request that the deadline for submitting proposals be extended at least until July 3. This would allow us to prepare properly and submit a maximally competitive proposal that meets all requirements.

We kindly ask you to consider this request and look forward to your comments.
	
Kaip jau buvome pažymėję 2026-06-19 atsakyme, Techninės specifikacijos 2.5 punkte nustatyta didžiausia leistina sieros vandenilio (H₂S) koncentracija dujose (≤ 0,007 g/m³) nustatyta vadovaujantis Lietuvos Respublikos energetikos ministro 2013 m. spalio 4 d. įsakymu Nr. 1-194 „Dėl Gamtinių dujų kokybės reikalavimų patvirtinimo“ (aktuali redakcija) (toliau – EnMin įsakymas). Šis EnMin įsakymu patvirtintas reikalavimas yra imperatyvus, ir negali būti  keičiamas.
Perkančiajam subjektui būtina, kad siūloma įranga atitiktų teisės aktais nustatytą reglamentavimą, būtų tinkama saugiai ir patikimai eksploatuoti esant visiems techninėje specifikacijoje nurodytiems dujų parametrams, įskaitant H₂S koncentraciją iki 0,007 g/m³.
Pažymime, kad, remiantis Specialiųjų pirkimo sąlygų (SPS) 7.2. punktu ir jo papunkčiais Tiekėjas kartu su paraiška turi pateikti visus šiame punkte ir jo papunkčiuose išvardintus dokumentus. Remiantis SPS 9.2. punktu ir jo papunkčiais kartu su pirminiu pasiūlymu tiekėjai turi pateikti visus šiame punkte ir jo papunkčiuose išvardintus dokumentus:
-Užpildytą ir kvalifikuotu elektroniniu ar fiziniu parašu pasirašytą Pasiūlymo formą (SPS 2 priedas).
- Jei Pasiūlymą elektroniniu ar fiziniu parašu pasirašo Tiekėjo vadovo įgaliotas asmuo, prie Pasiūlymo turi būti pridėtas galiojantis rašytinis įgaliojimas arba kitas dokumentas, suteikiantis teisę pasirašyti Pasiūlymą.
- Užpildyta Techninės specifikacijos priedas Nr. 1 , t. y. Naujo mobilaus dujų kompresoriaus techninių parametrų atitikties lentelė.
- Užpildytą Tiekėjo pasiūlymų dėl derėtinų sąlygų formą (SPS 12 priedas) (jei norima derėtis dėl sąlygų, nurodytų SPS 2.5. punkte).
- Dokumentai, nurodyti Techninės specifikacijos 3.16.12. punkte, įrodantys Techninės specifikacijos 3.16.11. punkto atitiktį.
Taip pat atkreipiame dėmesį, kad, remiantis TS 5 punktu ir jo papunkčiais, kartu su prekėmis tiekėjas turės pateikti išsamią techninę dokumentaciją ir  visus dokumentus, įrodančius ir patvirtinančius Prekių atitiktį techniniams reikalavimams. 



Dėl informacijos apie faktines H₂S koncentracijas gamtinių dujų perdavimo sistemoje pažymime, kad techninė specifikacija rengiama vadovaujantis galiojančiais teisės aktais ir projektavimo sąlygomis, todėl įrangos atitiktis turi būti užtikrinama visam leidžiamam dujų kokybės parametrų diapazonui. Faktinės istorinių laikotarpių dujų sudėties reikšmės nėra pagrindas nesivadovauti EnMin įsakyme nustatytomis H₂S reikšmėmis, todėl ši informacija nebus pateikiama.
Atsižvelgiant į tai, kad Paraiškų pateikimo terminas jau buvo pratęstas iki 2026 m. birželio 29 d., o visi pirkimo dokumentai, įskaitant Techninę specifikaciją, tiekėjams buvo prieinami nuo pirkimo paskelbimo dienos, Perkantysis subjektas vertina, jog tiekėjai turėjo pakankamai laiko susipažinti su reikalavimais ir pateikti klausimus dėl dokumentų paaiškinimo. Šiuo metu nustatytas Paraiškų pateikimo terminas yra proporcingas pirkimo objekto sudėtingumui ir laikomas pakankamu tinkamai parengti Paraiškas. Dėl šių priežasčių Paraiškų pateikimo terminas nebus pratęsiamas ir lieka galioti iki 2026 m. birželio 29 d.


 
	As previously noted in our response of 2026‑06‑19, the maximum permissible hydrogen sulfide (H₂S) concentration in gas (≤ 0.007 g/m³) specified in Clause 2.5 of the Technical Specification has been established in accordance with the Order No. 1‑194 of 4 October 2013 of the Minister of Energy of the Republic of Lithuania, “On the Approval of Natural Gas Quality Requirements” (as amended) (hereinafter – the EnMin Order). The requirement established by the EnMin Order is mandatory and cannot be amended.

The Contracting Entity requires that the proposed equipment comply with the regulatory framework established by legal acts and be suitable for safe and reliable operation under all gas parameters specified in the Technical Specification, including H₂S concentrations up to 0.007 g/m³.

We note that, pursuant to Clause 7.2 of the Special Procurement Conditions (SPC) and its sub‑clauses, the Supplier must submit all documents listed in this clause and its sub‑clauses together with the Application. Pursuant to Clause 9.2 of the SPC and its sub‑clauses, Suppliers must submit all documents listed in this clause and its sub‑clauses together with the Initial proposal:

A completed Proposal Form (Annex 2 to the SPC), signed with a qualified electronic or physical signature.

If the Proposal is signed by a person authorised by the Supplier’s manager, a valid written authorisation or other document granting the right to sign the Proposal must be attached.

A completed Annex No. 1 to the Technical Specification, i.e., the Compliance Table of Technical Parameters for the New Mobile Gas Compressor.

A completed Supplier’s Proposal on Negotiable Conditions Form (Annex 12 to the SPC) (if the supplier wishes to negotiate the conditions listed in SPC Clause 2.5).

The documents specified in Clause 3.16.12 of the Technical Specification proving compliance with Clause 3.16.11 of the Technical Specification.

We also draw attention to the fact that, pursuant to Clause 5 of the Technical Specification and its sub‑clauses, the supplier will be required to provide detailed technical documentation and all documents proving and confirming the compliance of the Goods with the technical requirements together with the delivery of the Goods.

Regarding information on actual H₂S concentrations in the natural gas transmission system, we note that the technical specification is prepared in accordance with applicable legal acts and design conditions; therefore, equipment compliance must be ensured for the entire permissible range of gas quality parameters. Actual historical gas composition values do not constitute grounds for disregarding the H₂S limits established in the EnMin Order; therefore, such information will not be provided.

Considering that the deadline for submitting Applications has already been extended until 29 June 2026, and that all procurement documents, including the Technical Specification, have been available to suppliers since the date of the procurement announcement, the Contracting Entity considers that suppliers have had sufficient time to familiarise themselves with the requirements and submit questions regarding clarification of the documents. The currently established deadline for submitting Applications is proportionate to the complexity of the procurement object and is deemed sufficient for proper preparation of Applications. For these reasons, the deadline for submitting Applications will not be extended and remains valid until 29 June 2026.


* Suinteresuoto (-ų) tiekėjo (-ų) prašymo (-ų) paaiškinti/ patikslinti Pirkimo dokumentus tekstas neredaguotas./ * The text of the request(s) by the interested supplier(s) to clarify or amend the Procurement Documents has not been edited.
** Paaiškinimas/ patikslinimas ir jo nuostatos turi viršenybę prieš ankstesnes Pirkimo dokumentuose išdėstytas nuostatas./ ** The clarification or amendment and its provisions take precedence over the previous provisions set forth in the Procurement Documents.
Pastaba: Paklausimų ir atsakymų  tekstas Lietuvių kalba turi viršenybę prieš vertimą į anglų kalbą. /Note: The Lithuanian text of the questions and answers takes precedence over the English translation. 



