Annex No. 3 to SPC

Radio Equipment Parameter Measuring Device

Technical Specification

This document describes requirements for Radio Equipment Parameter Measuring Device - Radio Communication Tester (hereinafter - RCT) for radio receivers and transmitters, which public limited liability company “Oro navigacija” intends to purchase.

- General

  - The RCT shall be suitable for testing analog RF signals of transmitters and receivers.

  - For receiver tests, audio signals from the internal generators shall be modulated onto the RF carrier.

  - The audio signals demodulated by the device under test shall be fed into the RCT inputs and then analysed.

  - For transmitter tests, the RCT shall demodulate the received signal and measure the demodulated audio signal and the RF signal.

  - The user-friendly interface shall be employed to allow users quickly reach all settings and easily perform measurements.

  - The measurement results shall be clearly and conveniently displayed.

  - The RCT shall be operable via the touchscreen. All functions shall be quickly accessible.

  - The rotary knob shall be available to change settings, to scroll through the frequencies and levels and immediately see the impact on the measurement results.

- Common requirements

  - The Supplier must be the manufacturer of the offered RCT or an authorised representative of the manufacturer entitled to sell the offered RCT and provide technical maintenance services for it, or must have concluded an agreement with another economic operator authorised by the manufacturer and entitled to provide technical maintenance services for the offered RCT during the warranty period.

  - The RCT shall have flexible internal switching capabilities for the audio and RF paths.

  - Users shall be able to configure the internal generators and measurements.

  - It shall be possible to set and invoke predefined test scenarios for receiver, transmitter and duplex tests, the RF and audio paths to save the time and eliminate configuration errors for standard test cases.

  - The RCT shall allow users to access all configuration options for expert measurements.

  - The RCT shall be suitable for VoIP generation and analyses and VoIP delay measurement (according to EUROCAE – ED137B/C).

  - The RCT shall be supplied with a calibration certificate issued by the manufacturer or an accredited calibration laboratory.

  - Calibration of RCT shall be according to ISO/IEC 17025 (or equivalent).

  - Delivery time: no later than 4 months from the date of contract signature. The RCT must be delivered to the following address: AB “Oro navigacija”, Balio Karvelio g. 25, LT-02184 Vilnius. The Supplier must ensure that the RCT is delivered on business days outside peak hours, in order to reduce the impact of delivery on traffic congestion, i.e. from Monday to Thursday from 10:00 to 16:30, and on Friday from 10:00 to 15:15. The Supplier must ensure that all additional documentation related to the RCT provided to the Buyer is submitted in electronic format. This requirement shall not apply to documents which are provided by the manufacturer together with the RCT in paper form as standard. Additional paper copies of documents shall be provided only upon the Buyer’s separate request.

- Radio frequency generator

  - Frequency range from 100 kHz to 3 GHz.

  - Frequency resolution 1 Hz.

  - COM output CW level range from -128 dBm to 2 dBm.

  - OUT output CW level range from -120 dBm to 10 dBm.

  - Output level resolution 0.01 dB.

  - Output level uncertainty less than 2 dB.

  - COM connector on front panel type N female, 50 Ω.

  - OUT connector on front panel type N female, 50 Ω.

  - VSWR less than 2.0.

  - Attenuation of second harmonics more than 30 dB.

  - Attenuation of third harmonics more than 40 dB.

  - Attenuation of non-harmonics more than 55 dB.

  - Modulation: AM, FM, PM.

  - AM depth from 0 to 100%.

  - FM deviation from 0 to 100 kHz.

  - PM deviation from 0 to 10 rad.

  - Modulation frequency from 0 Hz to 20 kHz.

  - Shall be protected from RF power overload.

- Radio spectrum analyzer

  - The spectrum analyzer shall have full span and user-defined span functionality.

  - The zero span mode shall enable analysis in the time domain.

  - Distorted signals shall be analyzed in the spectrum analyzer's time domain.

  - Depending on the type of modulation, either the modulation deviation or modulation depth shall be measured and displayed.

  - Analyzer shall have power and frequency metering functionalities.

  - Requirements for analyzer:

    - Frequency range from 100 kHz to 3 GHz,

    - Frequency resolution 1 Hz

    - COM input continuous power up to 100 W,

    - IN input (connector on front panel type N female, 50 Ω) up to 100 mW or 20 dBm,

    - OUT reverse input up to 20 mW or 13 dBm,

    - COM input power setting range from -13 to 50 dBm with attenuator, and from -30 to 36 dBm without attenuator,

    - IN input power setting range from -36 to 20 dBm.

    - VSWR less than 2.0.

    - Shall be protected from RF power overload.

- Audio frequency generators

  - The RCT shall be equipped with internal audio generators that can generate two tones simultaneously and modulate them onto the RF carrier.

  - The required levels shall be settable. If the signal is to be modulated onto the RF carrier, the modulation characteristics shall be configured.

  - Requirements for audio generators:

    - Output impedance less than 4 Ω,

    - AF1 OUT and AF2 OUT output connectors type BNC,

    - Frequency range from 20 Hz to 20 kHz,

    - Frequency resolution 1 Hz ,

    - Level range from 10 μV to 5 V,

    - DC component 0 V, must not be,

    - Level resolution 10 μV,

    - Output voltage uncertainty less than 1.5 V + amplitude resolution.

- Audio frequency analyzer

  - All audio signals both externally fed signals and demodulated audio signals shall be analyzed.

  - Highpass, lowpass and weighting filters shall be possible to apply to the audio signals.

  - The quality of the audio signal shall be determined with SINAD, S/N and THD, and shall be determined and displayed simultaneously.

  - The test frequency shall be selectable by the user.

  - Input impedance 100 kΩ, 15 pF.

  - AF1 IN and AF2 IN input connectors type BNC.

  - Frequency range from 50 Hz to 20 kHz.

  - Level range from 50 μV to 30 V.

  - Level resolution 1 μV at level up to 1 mV, and 0.1 % at level 1 mV and above.

  - SINAD and S/N measurement requirements:

    - Frequency range from 50 Hz to 10 kHz,

    - Measurement range from 1 to 46 dB,

    - Measurement level range (RMS) from 10 mV to 30 V,

    - Resolution 0.1 dB,

    - Uncertainty less than 1 dB at input level from 100 mV to 20 V.

  - THD measurement requirements:

    - Frequency range from 50 Hz to 10 kHz,

    - Level range (RMS) from 10 mV to 30 V,

    - Resolution 0.01 %,

    - Measurement uncertainty less than 1 dB,

    - Inherent distortion less than 0.05 % THD+N or less than 0.025 % THD, at input level from 100 mV to 20 V,

    - Uncertainty less than 2 % + inherent distortion.

- Adjacent channel power and occupied bandwidth

  - The ACP measurement shall be available of the power that a transmitter emits into adjacent channels.

  - Channel and measurement bandwidth settings shall be adjusted as needed from 100 Hz to 4 MHz.

  - Results shall be presented in graphical and tabular form.

  - The occupied bandwidth shall be measured to determine the bandwidths occupied by a settable percentage of the power.

- Oscilloscope

  - The oscilloscope shall show the audio signals of 20 kHz bandwidth that are fed into the AF1 IN and AF2 IN ports, including the demodulated audio signals for transmitter tests.

  - Marker functions shall simplify analysis of these audio signals.

  - Audio signals shall be viewed in both the time domain and in the frequency domain for easy and comprehensive analysis of all audio signals.

- Interferer

  - The RCT shall generate two RF signals. If these signals are positioned outside of the device receive window in such a way that at least one intermodulation product lies within the receive window, it is possible to assess the receiver quality.

  - The interferer shall allow users to measure co-channel rejection and adjacent channel suppression, eliminating the need to employ an additional generator to generate the interfering signal.

  - The RCT shall allow the user to generate the two RF signals at different levels within the available 20 MHz bandwidth. Both signals, the wanted signal and the interferer, shall be modulated independently of one another.

  - The levels of both signals shall be set independently of each other.

- Predefined test scenarios

  - Predefined scenarios for standard measurement tasks shall be available to enable users to configure the RCT software and hardware with a finger tap. Predefined scenarios shall be provided for TX measurements, RX measurements, spectrum analysis.

  - In expert mode, users shall be able to configure the RCT as required. Audio and RF paths shall be switched as needed. All generators and analyzers shall be accessible and configurable.

- Display of parameters and measurement results

  - Users shall have a choice of two modes for displaying parameters and measurement results:

    - The tab mode for displaying the values in detail. All generator and analyzer values shall be displayed in separate full-screen tabs.

    - The split screen mode shall be available for complete overview, where the generator and analyzer values are displayed simultaneously.

- Trim view

  - The trim view graphically shall display selected measurement values and their limits. In contrast to scalar displays, this view shall make it easier to recognize when the values fall below or exceed limits and facilitates comparison of transmitters and receivers.

- Remote control

  - Remote control via Ethernet/LAN interface shall be available for integration into automated test environments.

- Power supply

  - 230 V AC shall be used as the main power supply via Schuko type F plug.

  - DC power supply from the battery shall be available for field measurements.

  - Li-Ion battery and battery charger shall be added.

- Accessories

  - Hard case.

  - Soft case.

  - Set of measurement cables.

  - User manual.

  - Calibration certificate.

- Warranty

  - Warranty period shall be at least 24 months.

  - During this period the Supplier shall provide repair and calibration service at its own expense. If a specific calibration interval is not set out in the Technical Specification or in the manufacturer’s documents, calibration services shall be provided as reasonably necessary to ensure the proper functioning of the RCT and its compliance with the established requirements.

- Abbreviations

ACP – Adjacent channel power

RCT – radio communication tester

LAN – local area network

AM – amplitude modulation

FM – frequency modulation

PM – phase modulation

CW – continuous wave (non-modulated)

COM – combined (OUT and IN)

OUT – output

IN - input

AF – audio frequency

RF – radio frequency

TX – transmitter

RX – receiver

SINAD – signal-to-noise and distortion

S/N – signal-to-noise

THD – total harmonic distortion

N – noise

RMS - root mean square

VoIP – voice over internet protocol

VSWR – voltage standing wave ratio